Model SX712 ORP Meter

Operation Manual

Shanghai San-Xin Instrumentation Inc.

SX700 Series of Portable Meters

1. Model SX711 pH / mV Meter

2. Model SX712 ORP Meter

3. Model SX713 Conductivity / TDS / Salinity / Resistivity Meter

4. Model SX716 DO Meter

5. Model SX721 pH / ORP Meter

6. Model SX723 pH /mV / Conductivity Meter

7. Model SX725 pH / mV / DO Meter

8. Model SX726 Conductivity / DO Meter

9. Model SX731 pH / ORP / Conductivity Meter

10. Model SX736 pH / mV / Conductivity / DO Meter

11. Model SX751 pH / ORP / Conductivity / DO Meter

Table of Contents

1. Brief Introduction--------------------------------------------------------------------------------------

2. Technical Parameters ---------------------------------------------------------------------------------

3. Instructions to the Meter -----------------------------------------------------------------------------

3.1. LCD Display -------------------------------------------------------------------------------------

3.2. Operation Keys -----------------------------------------------------------------------------------

3.3. The Storage, Recall and Elimination of the Measuring Information ---------------------

4. ORP Measurement ------------------------------------------------------------------------------------

4.1. Solution Test --------------------------------------------------------------------------------------

4.2. Considerations -----------------------------------------------------------------------------------

4.3. Parameter Setting --------------------------------------------------------------------------------

5. Meter’s Complete Kit --------------------------------------------------------------------------------

6. Warranty ------------------------------------------------------------------------------------------------

1. Brief Introduction:

Thanks for buying and using the model SX712 ORP meter (the following called “meter” in short).

Before using this meter, please read the operation manual carefully in order to help use and maintain it correctly. On the basis of improving instrument of performance constantly, we reserve the right of changing the content of this manual and accessories in case of not notifying in advance.

This meter is a perfect combination with the most advanced electronic technology, sensor technology and software design. The meter can measure the parameters of ORP for high accuracy solution. It is the best portable ORP meter with the highest performance and the lowest cost. It is suitable for the trade such as the mining industry, water treatment projects and environmental protection, etc., especially has more extensive application in the field and spot test.

Built-in microprocessor chip, beautiful appearance, multi-functional and easy to use, this meter has the following prominent features:

1.1. Built-in microprocessor chip, with the intelligent functions of data storage, function setting and auto power off and low voltage display etc. Equips with ORP standard solution and special carrying case, easy to use.

1.2.	Adopts digital filter and step slipping technology to intelligently improve meter’s response speed and result accuracy. “     ” icon will appear when reading to be stable.

1.3.	Meter’s circuit board adopts SMT film-covering technology to improve meter’s production reliability.

1.4. Meter has the blue backlight LCD display.

1.5. Dustproof and waterproof meter meets IP57 rating.

2. Technical Parameters:

2.1. ORP:

| Measuring range | -1999mV to 0 to 1999mV |
| --- | --- |
| Resolution | 1mV |
| Accuracy | Meter：±0.1% FS, connect meter with electrode：±15mV |

2.2. Other Technical Parameters:

| Data storage | 100 groups |
| --- | --- |
| Storage content | Series number, measuring value and measuring unit temperature |
| Power | Two AA batteries (1.5V×2) |
| Size and weight | Meter: (65×120×31) mm/180g; Carrying case: (255×210×50) mm/790g |
| Quality and safety certification | ISO9001: 2000, CE and CMC |

2.3. Working Condition:

| Environment temperature | 5 to 35 ℃ (0.01 grade) |
| --- | --- |
| Environmental humidity humidity | ≤85% |
| IP rating | IP57 Dustproof and waterproof |

3. Instructions to the Meter:

3.1. LCD Display:

① —— Parameters mode icon

② —— Measuring value

③ —— Serial number and icon as measuring value to be stored and recalled; and the indication icons for the special measuring state.

M+ — measuring value to be stored icon;

RM — reading to be recalled icon;

④ —— Measuring unit

⑤ —— Measuring value to be stable icon

⑥ —— Indication icon of low voltage, it appears when the voltage less than 2.6V, call attention to change the batteries.

3.2. Operation Keys:

The meter has 5 operation keys in all.

3.2.1.         — Switch key

3.2.2.         — Calibration key

When in the parameter setting state, press this key to change the number.

3.2.3.         — Function key

Press key to enter into the parameter setting mode P1, P2.

3.2.4.         — Backlight and entrance key

(a) When in the measuring state, short press (less than 1.5s) to open or close the backlight display;

(b) When in the calibration state or the parameter setting state, press this key to confirm, and then the meter enters into measuring state;

3.2.5.         — The combination key of memory and recall

(a) Short time press (press time less than 1.5s) to save the measuring data； long time press (press time more than 2s) to recall the saved measuring data when in the measurement state.

(b) When in the parameter setting state, press this key to alter the number or ON/OFF state.

3.3. The Storage, Recall and Elimination of the Measuring Information:

3.3.1. Store the measuring information:

In the measuring mode, when the measuring data is stable and appear the “      ”, short-time press          key (less than 1.5s), LCD will display “ M+ ” icon and storage serial number, and meanwhile memory all the measuring information. Meter can separately store 100 groups of measuring information.

3.3.2. Recall measuring information:

(a) Under the measuring mode, long press the        key (press time more than 2s), meter will recall the last stored information, the storage number, “ RM ” icon and the complete measuring information will appear in the upper right corner of the LCD. Again press      or      key, meter will in turn recall all the measuring information, depress      or      key can quickly query measuring information under other serial number;

(b) In the recalling mode (there are “ RM ” and storage serial number in the upper right corner of the LCD), press         key to return to the measuring mode.

3.3.3. Eliminate the stored measuring information:

In the recalling mode, depress the         key for 5 seconds, LCD will appear “        ” for 2 seconds. It means the storage has been eliminated, and returns to measuring mode.

4. ORP Measurement:

4.1. Solution Test:

4.1.1. Press         key to turn on.

4.1.2. Connect the 301Pt-S ORP combined electrode, immerge it into the sample solution, slowly stir and then still. When there appear a “      ” and a stable reading, that is the ORP value.

4.2. Considerations:

4.2.1. When doing ORP measurement, the meter no need to be calibrated, but if you have problems about the quality of ORP electrode or the testing result, you can use the ORP standard solution to test the mV value of this meter, so you can check ORP electrode or meter if it is accuracy or not.

The ORP standard solution of meter equipment 222mV (25℃) ±15mV can keep for one year. Chart (4-1)

Chart (4-1)

| ℃ mV | ℃ mV | ℃ mV | ℃ mV |
| --- | --- | --- | --- |
| 10 242 | 20 227 | 30 215 | 38 205 |
| 15 235 | 25 222 | 35 209 | 40 201 |

4.2.2. The ORP electrode washed and activation：ORP electrode after being long time used, the polluting on the surface of platinum can lead the measure result is inaccurate and response slowly, so we following these steps to wash:

(a) To inorganic pollution, you can put electrode into 0.1mol/L dilute hydrochloric acid for 30 minutes after washing by the purified water, then put it into the electrode solution for 6h, then use it.

(b) To organic oil polluting and the oil film polluting, first can use abluent to wash the surface of the platinum, second wash it with the purified water, then put it into electrode solution for 6h, then use it.

(c) If the surface of the platinum was terribly polluted and became the oxidative film, first can use toothpaste to polish the surface of the platinum, second use the purified water to wash, then put it into electrode solution for 6h, then use it.

4.3. Parameter Setting:

4.3.1. ORP measuring parameter setting schedule (Chart (4-2))

(Chart (4-2))

| Prompt Mark | Parameter Setting Items | Code | Parameters |
| --- | --- | --- | --- |
| P1 | Backlight display time setting |  | 0-1-3-6 min |
| P2 | Auto power off setting |  | 0-10-20 min |

4.3.2. Backlight display time setting (P1)

(a) press the       key, the meter enters into mode P1, see        chart (4-1).

(b) Press         or         key to choose the time: 0min, 1min, 3min or 6min. The auto power off function will be closed if choosing 0min.

(c) Press        key to enter into next parameter setting or press         key to conform and return to measuring mode.

4.3.2. Auto power off setting (P2).

(a) Under P1 state, press         key to enter P2 mode, see chart (4-2).

(b) Press        or      key to choose the time: 0min, 10min or 20min. The auto power off function will be closed if choosing 0min.

(c) press         key to conform and return to measuring mode.

5. Meter’s Complete Kit:

5.1. Model SX712 ORP electrode

5.2. 301Pt-S plastic ORP combination electrode

5.3. 222mV ORP standard solution (50mL)

5.4. Spare AA batteries

5.5. Screw driver

5.6. Carrying case

6. Warranty:

6.1. We warrant this meter to be free of charge maintain, replace the parts or products under normal using circumstances, from purchased time within one year caused by manufacturing bad and unable to work.

6.2. Other attached electrodes do not belong to this warrant range except the temperature electrodes. But, if the newly purchased electrode went wrong without using, it’s free of charge to maintain and replace.

6.3. The above warranty is not apply to defects resulting from action of user such as misuse, improper wiring, operation outside of specification, improper maintenance or repair, or unauthorized modification.

Shanghai San-Xin Instrumentation, Inc.

Add: 3/F. Building No.4, No. 471 Guiping Road, Shanghai, China. 200233

Tel: +86-21-63362480                   Fax: +86-21-64956880

Website: www.shsan-xin.com            E-mail: wxmab@shsan-xin.com

# Text boxes

3

4

5

5

6

6

7

7

7

8

10

10

1 set

1 pc

1 btl

2 pcs

1 pc

1 pc